## Slide 1: Sensors & Actuators for Automatic Systems (S&AAS)
Dr. Imtiaz Hussain
Associate Professor
email: imtiaz.hussain@faculty.muet.edu.pk
URL :http://imtiazhussainkalwar.weebly.com/
Lecture-4
Mechanical Sensors

## Slide 2: Lecture Outline
Introduction
Potentiometric Sensors
Capacitive Sensors
Inductive Sensors
Applications

## Slide 3: Introduction
The class of sensors used for the measurement of mechanical phenomena is of special significance because of extensive use of these devices in process industry.
To learn to use the mechanical sensors it is essential to understand the mechanical phenomena themselves and the operating principle and application detail of sensor.

## Slide 4: Displacement, Location or Position Sensors
The measurement of Displacement, Location or Position is of prime importance in industries.
For example:
  Location and position of an object on conveyor system

## Slide 5: Displacement, Location or Position Sensors
The measurement of Displacement, Location or Position is of prime importance in industries.
For example:
  Orientation of steel plates in a rolling mill

## Slide 6: Displacement, Location or Position Sensors
The measurement of Displacement, Location or Position is of prime importance in industries.
For example:
  Liquid/ Solid level measurement

## Slide 7: Displacement, Location or Position Sensors
There are several types of sensors available:
  Potentiometric
  Capacitive
  Inductive
  Variable Reluctance

## Slide 8: Potentiometric Sensors
The simplest type of displacement sensor involves the action of displacement in moving the wiper of a potentiometer.

## Slide 9: Potentiometric Sensors
This device then converts linear or angular motion into a changing resistance that may be converted directly to voltage or current signal.
Such potentiometric devices often suffer from the obvious problem of mechanical wear, friction in the wiper action, limited resolution in wire-wound units, and high electronic noise.

## Slide 10: Potentiometric Sensors
Figure shows a simple potentiometric displacement sensor.
The resolution of this sensor is limited to the distance, , between the individual turns of wire, with the resulting resistance change of the single turn, .

## Slide 11: Potentiometric Sensors
Typical commercial units provide a choice of maximum shaft strokes from an inch or less to 5 feet or more.
Deviation from linearity of the resistance versus-distance specification can be as low as 0.1% to 1.0%.
Consider Fig. (b). If the circuit is unloaded, the output voltage V0 is a certain fraction of VT, depending on the position of the wiper:

## Slide 12: Example-1
A displacement transducer with a shaft stroke of 3.0 inch is applied in the circuit of Fig (b). The total resistance of the potentiometer is 5 k , and the applied voltage VT=5.0V. When the wiper is 0.9 inch from B, what is the value of the output voltage V0?

## Slide 13: Example-1
Solution:

## Slide 14: Example-2
A displacement transducer with a shaft stroke of 5.0 inch is applied in the circuit of Fig (b). The total resistance of the potentiometer is 10 k , and the applied voltage VT=5.0V. When the wiper is 1 inch from A, what is the value of the output voltage V0?

## Slide 15: Example-2
Solution:

## Slide 16: Potentiometric Sensors
To measure of the distance and direction that the object has traveled following arrangement is used.

## Slide 17: Example-3
A resistive position transducer with a resistance of 5000 and a shaft stroke of 5.0 inch is used in the arrangement of following Fig. Potentiometer R3 R4 is also 5000 , and VT = 5.0 V. The initial position to be used as a reference point is such that R1=R2 (i.e.. the shaft is at midstroke). At the start of the test, potentiometer R3R4 is adjusted so that the bridge is balanced (VE=0). Assuming that the object being monitored will move a maximum distance of 0.5 in. toward A, what will be the new value of VE?

## Slide 18: Example-3
At initial position
Voltage drop across R2 is
The value of R4 to balance the bridge (VE=0)

## Slide 19: Example-3
If the wiper moves 0.5 in toward A from midstroke, it will be 3.0 inch from B.
VE can now be calculated as:

## Slide 20: Rotary Potentiometer
The resistance between the wiper (slider) and "A" is labeled R1, the resistance between the wiper and "B" is labeled R2.
The total resistance between "A" and "B" is constant, R1+R2=Rtot.
If the potentiometer is turned to the extreme counterclockwise position such that the wiper is touching "A" we will call this θ=0; in this position R1=0 and R2=Rtot.
If the wiper is in the extreme clockwise position such that it is touching "B" we will call this θ=θmax; in this position R1=Rtot and R2=0.

## Slide 21: Rotary Potentiometer
R1 and R2 vary linearly with θ between the two extremes:

## Slide 22: Rotary Potentiometer
Potentiometer can be used to sense angular position, consider the circuit of figure-1.
Using the voltage divider principle we can write:
Figure-1

## Slide 23: Example-4
A potentiometric displacement sensor is to be used to measure work-piece motion from . The resistance changes linearly over this range from . Develop a signal conditioning circuit to provide a linear output of .
The key thing here is not to lose the linearity of the resistance versus displacement.
We cannot use a voltage divider here because voltage vs resistance variation is not linear for that circuit.
We can use an opamp circuit, because the gain and the output voltage is linearly dependent on the feedback resistor.
Solution:

## Slide 24: Example-4
We can use an opamp circuit, because the gain and the output voltage is linearly dependent on the feedback resistor.

## Slide 25: Capacitive Sensors
A second class of sensors for displacement measurement involves changes in the capacitance.
Basic operation of capacitive sensor can be seen from the familiar equation for a parallel-plate capacitor.
Where,

## Slide 26: Capacitive Sensors
From the equation it is clear that there are three ways to change the capacity.
  Variation of distance between plates (d)
  Variation of shared area of plates (A)
  Variation of dielectric constant (K)

## Slide 27: Example-5
Following figure shows the capacitive displacement sensor design to monitor small changes in work-piece position. Two metal cylinders are separated by plastic sheath of thickness and dielectric constant at of . If the radius is , (a) find the sensitivity in as the upper cylinder slides in and out of lower cylinder. (b) What is the range of capacity if varies from to ?

## Slide 28: Example-5
Capacity is calculated using following equation.
The net area is the area of the shared cylindrical area, which has a radius and height .
Thus , so the capacity can expressed as:

## Slide 29: Example-5
(a) The sensitivity with respect o the height, , is defined by how changes with ; that is .
Substituting the values, we get ()

## Slide 30: Example-5
(b) The minimum value of C can be calculated as.
Similarly Maximum value of C can be calculated as

## Slide 31: Inductive Sensors
If permeable core is inserted into an inductor, as shown in following figure, the net inductance is increased.
Every new position of the core produces a different inductance.
An ac bridge or other active electronic circuit sensitive to inductance may be employed for signal conditioning.

## Slide 32: Variable Reluctance Sensors
The variable reluctance displacement sensors differ from the inductive in that a moving core is used to vary the magnetic flux coupling between two or more coils, rather than changing an individual inductance.
Such device find applications in many circumstances for the measurement of both translational and rotational displacements.
Many configurations of this device exit, but the most common and extensively used is called Linear Variable Differential Transformer (LVDT).

## Slide 33: Linear variable differential transducer (LVDT)
LVDT linear position sensors are readily available that can measure movements as small as 0.001 mm up to 75 mm.

## Slide 34: Working principle of LVDT

## Slide 35: Working Principle of LVDT
k1 and k2 depend on the amount of coupling between the primary and the secondary coils, which is proportional to the position of the coil

## Slide 36: LVDT

## Slide 37: DC LVDT

## Slide 38: LVDT
As the core moves, these mutual inductances change, causing the voltages induced in the secondaries to change.
The coils are connected in reverse series, so that the output voltage is the difference (hence "differential") between the two secondary voltages.
When the core is in its central position, equidistant between the two secondaries, equal but opposite voltages are induced in these two coils, so the output voltage is zero.

## Slide 39: LVDT
When the core is displaced in one direction, the voltage in one coil increases as the other decreases, causing the output voltage to increase from zero to a maximum.
This voltage is in phase with the primary voltage. When the core moves in the other direction, the output voltage also increases from zero to a maximum, but its phase is opposite to that of the primary.
The magnitude of the output voltage is proportional to the distance moved by the core (up to its limit of travel), which is why the device is described as "linear". The phase of the voltage indicates the direction of the displacement.

## Slide 40: Example-4
An LVDT has a maximum, core motion of with the linearity of over that range. The transfer function is . If used to track the work-piece motion from to , what is the expected output voltage? What is the error in position determination due to nonlinearity?
Solution
Using the known transfer function, the output voltage can easily be calculated as:

## Slide 41: Example-4
Now
The linearity deviation shows up in the deviations of the transfer function. Thus the transfer function has the uncertainty of
This means that a measure voltage, , could be interpreted as a displacement that ranges from to which is approximately , as expected.

## Slide 42: Level Sensors
The Measurement of solid or liquid Level calls for a special class of displacement sensors.
The level measurement is most commonly associated with material in tank or hopper.
A great variety of measurement techniques exist, which are discussed here.

## Slide 43: Mechanical Level Sensor
One of the most common technique for level measurement, particularly for liquid, is a float that is allowed to ride up and down with level changes.
This float is connected by linkages to a secondary displacement measurement system such as potentiometer or LVDT.

## Slide 44: Electrical Level Sensor
There are several purely electrical methods for measuring level.
For example, one may used inherent conductivity of a liquid or solid to vary resistance seen by probes inserted into material.
Another common technique is shown here

## Slide 45: Example-5
The level of ethyl alcohol is to be measured from 0 to 5m using a capacitive system. Following specifications define the system.
  For ethyl Alcohol K=26 (for air K=1)
  Cylinder separation: d=0.5cm
  Plate area:
  Where R=5.75cm (average radius)
  L=distance along cylinder axis
Find the capacity variation as the alcohol level varies from 0 to 5m.

## Slide 46: Example-5
Solution
We all need to do is find the capacity for entire cylinder with no alcohol and then multiply it with 26.
Thu for air,
With the ethyl alcohol,

## Slide 47: Technologies
Through Air
Radar
Guided Wave
Radar
Ultrasonic
LEVEL Measurement Through Radar and Ultrasonic

## Slide 48: How it works
  The time it takes for the instrument’s signal to leave the antenna, travel to the product, and return to the antenna is calculated into distance.
  The instrument is spanned according to the distance the 100% and 0% points within the vessel are from its reference point.
  The measured distance can then be converted into the end user’s desired engineering unit and viewed on the head of the instrument or remote display.
100%
0%

## Slide 49
Through Air Radar

## Slide 50: Radar Technology – How it works
Radar is a time of flight measurement.
  Microwave energy is transmitted by the radar.
  The microwave energy is reflected off the product surface
  The radar sensor receives the microwave energy.
  The time from transmitting to receiving the microwave energy is measured.
  The time is converted to a distance measurement and then eventually a level.

## Slide 51: Radar Technology – Why use it?
Radar level measurement
Top mounted
Solids and liquids applications
Non-contact
RADAR is virtually unaffected by the following process conditions:
  Temperature
  Pressure and Vacuum
  Conductivity
  Dielectric Constant (dK)
  Specific Gravity
  Vapor, Steam, Dust or Air Movement

## Slide 52
Guided Wave
Radar
(TDR)

## Slide 53: Principle of Operation
A microwave pulse (2 GHz) is guided along a cable or rod in a 20” diameter or inside a coaxial system.
The pulse is then reflected from the solid or liquid, back to the head of the unit.
The travel time of the pulse is measured and then converted to distance.

## Slide 54: Guided Wave Radar Measurement
Guided Wave Radar level measurement
Time of Flight
Top mounted
Solids and liquids applications
Contact Measurement
GUIDED WAVE RADAR is virtually unaffected by the following process conditions:
  Temperature
  Pressure and Vacuum
  Conductivity
  Dielectric Constant (dK)
  Specific Gravity
  Vapor, Steam, or Dust Air Movement
  Build up (depends on type of build up)
  Foam

## Slide 55
Ultrasonic

## Slide 56: Ultrasonic Level Measurement – How it works
Time of Flight Technology
Short ultrasonic impulses emitted from transducer
Bursts are created from electrical energy applied to piezeo electric crystal inside the transducer
The transducer creates sound waves (mechanical energy)
With longer measuring ranges a lower frequency and higher amplitude are needed to produce sound waves that can travel farther
The longer the measuring range the larger the transducer must be

## Slide 57: Ultrasonic Level Measurement
Ultrasonic level measurement
Time of Flight
Top mounted
Solids and liquids applications
Non-contact
ULTRASONIC is virtually unaffected by the following process conditions:
  Change is product density (spg)
  Change in dielectric constant (dk)

## Slide 58: Ultrasonic Level Technology – Advantages
Can be mounted in plastic stilling wells
Narrow beam angles minimize effect of obstructions
Swivel flange available for applications with angles of repose
Familiar technology throughout the industry, therefore, often a trusted technology throughout the industry
Cost-effective

## Slide 59: End of Lecture-4
To download this lecture visit
http://imtiazhussainkalwar.weebly.com/